## Slide 1
Consiglio Nazionale delle Ricerche
Istituto di Fotonica e Nanotecnologie - Milan, Italy
NIREOS S.R.L. – Official Spin-Off company of Politecnico di Milano
Milan, Italy
Idea I-2021-03511 - TWINS-HG
An Ultracompact Hyperspectral imager in the Thermal Infrared
BBT Materials Processing, Ltd., Prague, Czech Republic
ESA CONTRACT NO. 4000137531/22/NL/GLC/OV
Final Presentation

## Slide 2: Project outline

## Slide 3: WP1: case study result
Visible / NIR
scattered solar radiation
spectral information about absorption, reflection and scattering of molecules and particles
land, water and atmosphere environments;
MWIR / LWIR
thermal radiation from Earth surface and atmosphere
spectrum reveals temperature, albedo, radiance and emissivity, composition of the surface
Bands due to absorption lines from vibrational modes
fingerprinting of chemical compounds on surface and in atmosphere.
3-14 μm: Thermal infrared (TIR)

## Slide 4: WP1: case study result
Volcanoes and Earthquakes
transient thermal anomalies preceding eruptions
Atmospheric gases (SO2, ash and water ice in the eruptive plumes)
Wildfires
burning biomass
Water Use and Availability
global freshwater supplies
Water resources
Ice
Earth Surface Composition and Change:
composition and thermal properties of the surface of the Earth
Spectral imaging of Earth surface and atmosphere:
enables monitoring various ecosystem and natural aspects, such as

## Slide 5: WP2: definition of instrument requirement
Requirement | Threshold | Target | Obtained
Spectral range | 7 – 15 um | 2 – 20 um | 3 – 16 um
Spectral resolution | 20cm-1 | 5cm-1 | 4.2 cm-1
Footprint | < 150 x 150 mm (225 cm2) | < 100 x 100 mm (100 cm2) | 120 x 90mm (108 cm2)
Interferogram Contrast | 50% | 70% | > 91.8%
Following the results from the case study, the following requirements have been defined:

## Slide 6: WP3: instrument design
Image 2D
Requirements?
Accuracy of delay < 1/100 of the optical cycle (2·10-15 seconds @ 600 nm)
HYPERspectral Image
Ray of light
0D
Interferometry + Fourier transform
- 1822 -
Wiener-Khinchin-Einstein theorem
- 1930 – 1934 – 1914 -
Fourier-transform spectroscopy

## Slide 7: CPI: A Common-Path Interferometer
z
x
y
45°
45°
da
db>da

D. Brida, C. Manzoni, G. Cerullo, Opt. Lett. 37, 3027 (2012)
C. Manzoni, D. Brida, G. Cerullo, US Patent: 9182284 (2015)
Generation of phase-locked replicas by birefringence
ordinary and extraordinary polarizations: different propagation speeds
Total delay: proportional only to thickness
Collinear replicas
Crystal Range (micron)
Alpha-BBO 0.2-3.5
Yttrium Vanadate 0.4-5
Lithium Niobate 0.4-5.2
Calomel 0.4-20

## Slide 8: CPI: A Common-Path Interferometer
x
y
45°
45°
da
db(x)
z

x
POL. 45°
D. Brida, C. Manzoni, G. Cerullo, Opt. Lett. 37, 3027 (2012)
C. Manzoni, D. Brida, G. Cerullo, US Patent: 9182284 (2015)

## Slide 9: CPI characterization - Dynamic Reproducibility
Interferograms acquired after 30 minutes
perfect reproducibility
x
t
PHOTO-DIODE
Delay or Motor position
D. Brida, C. Manzoni, G. Cerullo, Opt. Lett. 37, 3027 (2012)
Dynamic Reproducibility
SUPERCONTINUUM
FIBER LASER
Broadband incoherent
light source
BBO

## Slide 10: CPI characterization - Static Stability
delay fluctuations:
λ / 360 (∼5 as at 600 nm)
λ / 225 (∼3.87 as at 261 nm)
@ Fixed position of the wedge
x
DETECTION WAVELENGTH
t
Static Stability
D. Brida, C. Manzoni, G. Cerullo, Opt. Lett. 37, 3027 (2012)
SUPERCONTINUUM
FIBER LASER
Broadband incoherent
light source
0 50 100
Time [Minutes]
Phase [rad]
-0.1 0 0.1
Wavelength [nm]
258 261 274
TWINS characterization for experiment at FERMI
BBO

## Slide 11: WP4: Hyperspectral Imager: breadboarding and testing
Compact footprint
108 cm2
Lightweight
Optics+mechanics <1 kg
Robust
No stability control is required
Power consumption
< 50 W
Optical alignement
Alignment free
Ideal device for portable, on field and spaceborne applications
Optics and mechanics: to be qualified for space
4-5 cm
Insensitive to any mechanical fluctuations
Keeps alignment even in the presence of considerable mechanical perturbations

## Slide 12: Telescope design including the CPI
Sensor: Sony Pregius IMX174
Pixel size: 5.86 μm
N of pixels: 1280 x 1024
Horizontal AFOV=0.73°
Vertical AFOV=0.52°
Objective lens= 150 mm achromatic doublet
N-BK7 / N-SF5
Eyepiece lens= 50 mm achromatic doublet
N-BK7 / N-SF5
Camera lens= 150 mm achromatic doublet
N-BK7 / N-SF5
Beams are NOT collimated at the CPI plane
The interference contrast is good
The beams are parallel at the image plane (Image tele-centricity)
The beams corresponding to off-axis points are parallel at the CPI plane
  The interferogram phase is the same for each point of the object, leading to hyperspectral images without spatial aberrations
  Ideal for satellite staring images
More versatile: the camera lens can be changed based on the required AFOV, without affecting the interferential pattern spatial distribution.
Notes: Hyperspectral Cameras are very complex and expensive devices

## Slide 13: Picture of the Brassboard
Notes: Hyperspectral Cameras are very complex and expensive devices

## Slide 14: Long Distance Measurement
RGB projection of the hyperspectral data
Object: “Pizzo dei Tre Signori” (2554 m s.l.m)
Distance: 60 km
Horizontal AFOV=0.73°
Vertical AFOV=0.52°
Shot with mobile phone
Notes: Hyperspectral Cameras are very complex and expensive devices

## Slide 15: Test images at selected bands
450 nm
550 nm
650 nm
750 nm
850 nm
950 nm
Notes: Hyperspectral Cameras are very complex and expensive devices

## Slide 16: Interferometer test plan
Test plan | Test procedure | Spectral Range | 
Transmission of the interferometer | Measure the ratio between input and output powers | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
Interferometric contrast with collimated beam | Measure the interferogram as a function of the delay | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
Notes: What is the size of the laser beam in the TIR?
The shock test

## Slide 17: Tests in Visible: coherent light
VISIBLE RANGE
Notes: Hyperspectral Cameras are very complex and expensive devices

## Slide 18: Transmission and contrast
Reflectivity
4 calomel wedges: each beam experiences 4 reflections as ordinary beam and 4 reflections as extraordinary beam
  Expected from theory (Fresnel reflections: Rp=20.08%, Rs=10.58%): 24%
  Measured: 14 % Most light is scattered
VISIBLE RANGE
Delay [fs]
He:Ne laser @ 632 nm
Contrast
Contrast: 83 %

## Slide 19: Contrast in the visible
Transmitted light, on a screen
VISIBLE RANGE
Notes: Hyperspectral Cameras are very complex and expensive devices

## Slide 20: Visual Tests in Visible: incoherent light
Test with a Light bulb
Hypebolic spatial intereference pattern
Very limited spatial walk-off
Very high coherence
VISIBLE RANGE

## Slide 21: Characterization of interefrometer performances
  Transmission
  Interferometric contrast
  spectral resolution
  frequency calibration
Characterization of the following properties:
Notes: What is the size of the laser beam in the TIR?
The shock test

## Slide 22: Test plan
Test plan | Test procedure | Spectral Range | 
Transmission of the interferometer | Measure the ratio between input and output powers | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
Interferometric contrast with collimated beam | Measure the interferogram as a function of the delay | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
14%
82%
Notes: What is the size of the laser beam in the TIR?
The shock test

## Slide 23: Test Infrared: Goals & Methods
Characterize of the Calomel interferometer in terms of:
  Transmission
  Interferometric contrast
  spectral resolution
  frequency calibration
The light source
We used two light sources from COSMOS lab at Politecnico Lecco (Prof. Davide Gatti):
Narrowband: Quantum Cascade Lasers (QCL). The available QCL can be tuned from 5 to 9 micron
Broadband: difference frequency generation (DFG)
QCL lasers can access the IR range by exploiting multiple intersubband transitions in the conduction band of a structure made by repeated cells of 3 quantum wells. These are usually made of InGaAs, and separated by AlInAs spacers to create a miniband where electrons are injected or extracted.
TIR RANGE

## Slide 24: The experimental Setup
Gold mirror
Gold mirror
Gold mirror
Gold mirror
Gold mirror
Calomel
interferometer
Iris or Polarizer 1
Iris or Polarizer 2
Bristol wave-
meter
Laser
Chopper
Piezoelectric motor
MCT detector
Chopper driver
Signal
Reference
Lock-in
amplifier
Motor driver
PM 1
PM 2
PC
QCL
Gold mirror 2
Gold mirror 1
Chopper
Gold mirror 3
Iris or Polarizer 1
Iris or Polarizer 2
MCT Detector
Bristol wavemeter
Calomel
TIR RANGE

## Slide 25: The Detectors - I
Mercury-Cadmium-Telluride (MCT): Code: PCI-3TE-12-1×1-TO8-wZnSeAR-36 (probably) Type: epitaxial HgCdTe heterostructure photoconductive S/N: 443150140
Range: 1-14 μm (optimized for 12 μm) Window: wedged zinc-selenide anti-reflection coated
Mounted on MIP-DC-5M-F-M4 preamplifier
Placed after the interferometer, it collects the interferogram
TIR RANGE

## Slide 26: The detectors - I
Wavemeter/analyser (Bristol 771): The wavemeter analyser is a calibrated commercial device to measure the emission spectra from the QCL and the laser sources.
Type: Michelson interferometer, FT
Range: 0.375-12 μm
Resolution: 0.1 pm Accuracy: 0.2 ppm
PM16-401 Type: thermopile S/N: 1905780
Range: 0.19-20 μm Maximum power: 1 W
Power meter (PM):
It is used to measure the power of the laser light, for the characterization of the transmission
TIR RANGE

## Slide 27: Transmission @5 μm
Power measurements:
Light: QCL,
Input power: (PM 1, no polarizer)
Output power:
  w/o output polarizer:
  (PM 2)
  w output polarizer (max: transmission is optimized): (PM 2)
Transmission:
(predicted value: 20%) (predicted value: 4.4 %)
Date: 21/04/2023
Transmission is higher than expected because light polarization is p (lower Fresnel losses).
Method:
Transmission is calculated by taking two transmission measurements: Input: before interferometer; output: after interferometer
TIR RANGE

## Slide 28: Test plan
Test plan | Test procedure | Spectral Range | 
Transmission of the interferometer | Measure the ratio between input and output powers | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
Interferometric contrast with collimated beam | Measure the interferogram as a function of the delay | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
14%
82%
9.5 – 29.3%
Notes: What is the size of the laser beam in the TIR?
The shock test

## Slide 29: Contrast @5 μm
Measurement parameters
Laser: QCL, Lock-in Integration time Motor scan: 0 → 0.5 mm Number of steps: 201 Step size: 2 µm Step waiting time: 200 ms
Interferogram 1 is saturated
Date: 21/04/2023
Experimental Interferogram
Contrast
Method:
Contrast is measured by a single, short scan, and comparing maxima and minima
TIR RANGE

## Slide 30: Contrast @7 μm
Measurement parameters
Laser: Daylight QCL, (not measured) (input power)
Lock-in Integration time Motor scan: 0 → 0.5 mm Number of steps: 201 Step size: 2 µm Step waiting time: 200 ms
Date: 20/04/2023
Experimental Interferogram
Contrast
TIR RANGE

## Slide 31: Contrast @7 μm
  Measurement parameters: Laser: Daylight QCL, (measured with Bristol wavemeter) (PM 1, no polarizer) (PM2, no polarizer) Lock-in Integration time Motor scan: 0 → 0.5 mm Number of steps: 201 Step size: 2 µm Step waiting time: 200 ms
  Measured Transmission
  (in agreement with previous ones).
Date: 21/04/2023
Contrast
TIR RANGE

## Slide 32: Contrast @9 μm
Measurement parameters: Laser: QCL, (measured with Bristol) (PM before TWINS) (PM 2)
Lock-in Integration time Motor scan: 0 → 0.5 mm Number of steps: 251 Step size: 2 µm Step waiting time: 200 ms
Date: 27/04/2023
Contrast
TIR RANGE

## Slide 33: Test plan
Test plan | Test procedure | Spectral Range | 
Transmission of the interferometer | Measure the ratio between input and output powers | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
Interferometric contrast with collimated beam | Measure the interferogram as a function of the delay | CW laser beam in the visible (He:Ne) | CW laser beam in the TIR (Quantum cascade laser @5.1 micrometers)
14%
82%
9.5 – 29.3%
>92%
Notes: What is the size of the laser beam in the TIR?
The shock test

## Slide 34: Spectral Resolution @7 μm
Measurement parameters Laser: Daylight QCL, (measured with Bristol) (PM 1, no polarizer) , no polarizer)
Motor scan: -5 → 17 mm Number of steps: 7333 Step size: 3.0 µm
Date: 21/04/2023
Spectrum
after FT
Longest available scan,
Higher spectral resolution
These artefacts arise from errors in motor positioning
Motor position calibration is required

## Slide 35: Delay Calibration
Motor delay correction: obtained from time traces of the CW lasers at 5, 7 and 9 μm. The correction (stable and reproducible) enables to calculate the correct motor position, before taking the FT. All corrections show a similar behaviour, with only an offset. Correction is of the order of +/-2 um. Origin: unknown.
Date: 27/04/2023

## Slide 36: Frequency Calibration
From the spectra of the QCL laser we derived the frequency calibration of the interferometer. This relates the pseudofrequency to the optical frequency (measured with Bristol).
Date: 27/04/2023
FT Spectra after frequency and delay calibration
5.263 μm
1900 cm-1
FWHM:
4.3 cm-1
11.9 nm
7.675 μm
1303 cm-1
FWHM:
4.23 cm-1
24.9 nm
9.073 μm
1102 cm-1
FWHM:
4.2 cm-1
34.6 nm
5,5

## Slide 37: Contrast and spectral resolution
Wavelength | Contrast | Spectral width
5.3 μm | 96.1% | 4.30 cm-1
7.7 μm | 92.8% | 4.23 cm-1
9.1 μm | 97.6% | 4.20 cm-1
From the spectra of the QCL laser we derived the spectral resolution of the interferometer/spectrometer.

## Slide 38: Tests with Broadband Laser Spectrum
Date: 27/04/2023
Measurement parameters
Lock-in Integration time Motor scan: -3 → 3 mm Number of steps: 1500 Step size: 4 µm Step waiting time: 250 ms
Data treatment:
interferograms are apodized
Delays are corrected
spectra are frequency-calibrated
Source
Ultrabroadband laser source from difference frequency generation (DFG)
Wavelength:

## Slide 39: Broadband Laser Spectrum
Measurement 1
Date: 27/04/2023
Measurement 2
Laser: DFG,
Laser: DFG, + bandpass filter @
=
before TWINS
Spectra
after FT
Comparison between calomel interferometer and Bristol commercial spectrometer
Interferograms

## Slide 40: Broadband Incoherent Light
Chopper
Chopper driver
Calomel
MCT
Polarizer
Iris
Polarizer
Iris
Source:
Kantal (FeCrAl) filament
Parabolic mirror
Flat mirror
Flat mirror
Lock-in amplifier
Shield
Date: 28/04/2023
The ultimate test for the project: proving that the interferometer works also with incoherent radiation

## Slide 41: Test with Incoherent Light
Date: 28/04/2023
Contrast
Spectrum
We collected three scans with different scan ranges (influencing spectral resolution)
Interferogram
Spectral range:
3 - 16 μm

## Slide 42: Tests with microbolometer + incoherent source
Ceres V640 GigE
Dione 640 OEM thermal imaging core 640x480 pixels, 12 μm pixel pitch
Spectral range: 8 - 13 μm

## Slide 43: Thermal emission from heated samples
Spectrum
Interferogram
3 samples: pure monocrystalline quartz, defected mineral quartz, polycrystalline fused silica glass
the temperature: ≈200°C (473.15K), to increase the emittance (Stefan-Boltzmann law)

## Slide 44: Verification Matrix
Requirement | Threshold | Target | Obtained | 
Spectral range | 7 – 15 um | 2 – 20 um | 3 – 16 um | Threshold
Spectral resolution | 20cm-1 | 5cm-1 | 4.2 cm-1 | Target
Footprint | < 150 x 150 mm (225 cm2) | < 100 x 100 mm (100 cm2) | 120 x 90mm (108 cm2) | Target
Interferogram Contrast | 50% | 70% | > 91.8% | Target
The spectral range is only limited by emission of the light sources and the detection range
The spectral resolution has been reached by performing the longest scan, with the new interferometer design
The contrast has been greatly improved by the new interferometer design

## Slide 45: Communications to conferences
Small Satellites Systems and Services: 4S Symposium
16-20th May 2022 Vilamoura (Portugal)
F. Preda, D. Polli, G. Cerullo, C. Manzoni, “A compact and rugged Hyperspectral Camera for Remote Sensing based on a birefringent ultrastable common-path interferometer” (Presenter: C. Manzoni)
14 th IAA Symposium on Small Satellites for Earth System Observation
7-12th May 2023, Berlin (Germany)
M. Corti, B. Ardini, F. Preda, A. Perri, D. Polli, G. Cerullo, C. Manzoni, “A compact and rugged hyperspectral camera for Remote Sensing based on FT spectrometry” (Presenter: M. Corti)

## Slide 46: Future development plan
The further development will consist on the following activities:
optimization of the interferometer throughput: design of a suitable antireflection coating; optimization of the deposition procedures
engineering of the detection system: selection of a suitable focal plane array (micro-bolometer or MCT)
design and realisation of an imaging telescope
design of the calibration unit
Elegant Breadboard realisation and validation of the overall concept
The results of the first two activities will enable the development of a commercial product for terrestrial applications by NIREOS. A potential market for the TIR has already been identified.

## Slide 47
Consiglio Nazionale delle Ricerche
Istituto di Fotonica e Nanotecnologie - Milan, Italy
NIREOS S.R.L. – Official Spin-Off company of Politecnico di Milano
Milan, Italy
Idea I-2021-03511 - TWINS-HG
An Ultracompact Hyperspectral imager in the Thermal Infrared
BBT Materials Processing, Ltd., Prague, Czech Republic
ESA CONTRACT NO. 4000137531/22/NL/GLC/OV
Final Presentation