## Slide 1
Idea I-2023-01289
Real-time in-situ monitoring of particle deposition with on-chip electronic dust counter (DUST : DUst System Tracker)
Final presentation slides

## Slide 2: Synoptics of DUst System Tracker project

## Slide 3
DUST sensor system realization
Power supply : +5V (USB) or 3.6<Vext<6V, Imax=100mA (External) or Power bank (trickle charge) supporting Imin=37mA
Total Power consumption : 55 mW
A two boards system

## Slide 4
27mm
30.5mm
Micro-computer board
STm micro controller
USB
MicroSD
External power supply
Sensor chip connector
Data transfer & firmware upload : USB, Bluetooth, UART, STm link
Bluetooth

## Slide 5
18 mm
20 mm
Sensor-chip board
Test pads
C board connector
Sensor chip
AMS technology, 0.35um
2 mm

## Slide 6
Principle of operation
Interdigitated electrodes are designed in a differential scheme to isolate PM signal variation from sensor overall impedance value
Particle Matter (PM) changes the impedance between metal electrodes
Electrodes are stimulated at fixed frequency (Ex. 100kHz) and the signal is extracted with a lock-in amplifier operating at the same frequency
M.Carminati, G.Ferrari, F.Guagliardo, M.Sampietro “ZeptoFarad capacitance detection with a miniaturized CMOS current front-end for nanoscale sensors”, Sensors and Actuators A, 172, pp.117-123 (2011)

## Slide 7
Chip layout & functionalities
Digital to Capacitance Converter (DCC)
Interdigitated Electrodes
Low-noise
amplifiers
Lock-in & Gm-C Low-pass filter
2 mm
32 Channels
Sensing area (green area): 0.5 mm x 2 mm, subdivided in 32 channels
Interdigitated electrodes gap =1 μm
VAC = 3.3V
Low pass filter of the Lock-in amplifier, BW = 100 Hz
Recordings : number of particles/ch pulse amplitude time of arrival

## Slide 8
Architecture internal to the chip
Chip output is a differential DC voltage proportional to the measured PM event
Circuit to reset the sensor when many particles have been depositing on the electrodes
P. Ciccarella, M. Carminati, M. Sampietro, G. Ferrari, “Multichannel 65 zF rms Resolution CMOS Monolithic Capacitive Sensor for Counting Single Micrometer-Sized Airborne Particles on Chip,” IEEE J. Solid-State Circuits, Vol. 51, No. 11, pp. 2545–2553, (2016)

## Slide 9: Monitoring & Acquisition of Sensor outputs
Sensor CHIP
Two 12-bit ADCs internal to the STm32
TI ADS8867
16-bit
Differential reading
DC monitoring
for particle counts & tracking
Ch1
Ch2

for channels monitoring

## Slide 10: Monitoring & Acquisition of Sensor outputs
Sensor CHIP
Full scale 3.3V (255)
Mid value 1.65V (128)

## Slide 11
100
550
Detection of plastic beads 8m (1)
Thermo Scientific Fluoro-Max (red fluorescent polymer microspheres)
Particles powder 8um in diameter
F.Zanetto, F.Ciobanu, M.Carminati, G.Ferrari, M.Sampietro, J.Eck “Portable fallout monitoring instrument for particle tracking in space applications”, ISMSE-ICPMSE, Saint Raphael (France) 7-11 Oct. 2024

## Slide 12
Detection of plastic beads 8m (2)

## Slide 13: Dust particles counting
Example of outcome when a blow of microbeads of 8um in diameter are depositing on the surface of the detector
If the platform is not connected to the computer, PM is still counted (internal memory of STm or uSD) and once it is reconnected, all the data (number of particles, time instant and signal amplitude) are transmitted to the computer.

## Slide 14
Noise performance on 32 channels
All channels are very low noise, ensuring very high resolution in the PM detection (signal of 8um bead is about 600 bins, channel noise is <10 bins)

## Slide 15
Granulometric Spectrum
Detectable particle size: between 1 μm and about 20 μm
A different range can be targeted with a different geometry of electrodes. Example: with electrodes gap of 5μm, the range would be 5μm<PM<100 μm.

## Slide 16
Data handling & communication method
b) Data memorized in the internal memory of STm C
. number of particles detected in each channel . transmission via Bluetooth
Data can be acquired, collected and transmitted in many ways :
c) Data memorized in the SD during experiment and transmitted off-line with USB/UART/Bluetooth
+ sampled values around the event
+ time of arrival of particles
a) USB to computer in run time
+ raw sampling data on 32 ch.
+ time of arrival of particles + circuit and sensor operational parameters
+ access to C

## Slide 17: Thermal behaviour of the system
The total Power consumption of 55 mW localizes in the sensor chip and in the 5V-to-3.3V DC-DC converter

## Slide 18: Insensitivity to ambient light
Light impinging onto the sensor produces an output signal of max 60 bins, to be compared with the 8um PM signal of 600 bins.
OFF
Light ON
OFF
ROOM light
Light ON
OFF
OFF
SPOT light

## Slide 19: Dissemination
Conference partecipation

## Slide 20
Gantt chart
20

## Slide 21: The POLIMI staff of DUST project
Florin Ciobanu
Adjoint researcher
Francesco Zanetto
Research Associate
Giorgio Ferrari
Associate Professor
Marco Carminati
Associate Professor
Marco Sampietro
Full Professor
Innovative Integrated
Instrumentation for Nanoscience
I3N